## Slide 1
The JenaValve Trilogy™ Heart Valve System in High Surgical Risk Patients with Symptomatic, Severe Aortic Regurgitation
Torsten P. Vahl, MD
and Vinod H. Thourani, MD for the ALIGN-AR Investigators
2-year Outcomes From The ALIGN AR Trial
The Trilogy THV System is for Investigational Use Only in the United States and is Limited by Federal (or United States) law for this use.

## Slide 2
Within the prior 24 months, I have had a relevant financial relationship with a company producing, marketing, selling, re-selling, or distributing healthcare products used by or on patients:
Nature of Financial Relationship | Ineligible Company
Grant/Research Support | Edwards Lifesciences, Boston Scientific, Medtronic, JenaValve
Consultant Fees/Honoraria | 4C Medical, Abbott, Innovheart, Philips
Individual Stock(s)/Stock Options | None
Royalties/Patent Beneficiary | None
Executive Role/Ownership Interest | None
Other Financial Benefit | None
Disclosure of Relevant Financial Relationships
All Relevant Financial Relationships have been mitigated.
Faculty disclosure information can be found on the app

## Slide 3: Purpose
Report 2-year outcomes from the ALIGN-AR pivotal trial evaluating TAVR with the Trilogy transcatheter heart valve (THV) amongst high surgical risk patients with symptomatic ≥3+ aortic regurgitation

## Slide 4: Trilogy Porcine Pericardial Valve
Aligns THV with native cusps
Locators “clip” onto native leaflets forming a natural seal and stable securement
Large open cells provide access to low coronaries
Flared sealing ring conforms to annulus
27-31F opening
Alignment
Positioning/Anchoring
Deployment

## Slide 5: ALIGN AR Study Design
Multicenter, Non-blinded, Single Arm Evaluation of Patients with Symptomatic ≥3+ Aortic Regurgitation at High Risk for Surgery
Trilogy THV Implantation
Clinical Evaluation, Echocardiography, Functional and QoL Assessment at 30 Days, 6 Months, 1 Year and Annually up to 5 Years
30 Day Primary Safety Endpoint
1 Year Primary Efficacy Endpoint
Comparison with Prespecified Performance Goal

## Slide 6: Consort Diagram
180 Patients Enrolled in ALIGN-AR
180 (100%) included in 1-year follow up
Lost to Follow-up (n=1)
Day 561
Withdrew (n=2)
Day 408
Day 440
176 (98%) Included in 2-year follow-up
Valve Explant for Endocarditis (n=1)
Day 380

## Slide 7
Chart: 
| 0.1 | 0.2 | 0.3
Series 1 | 0 | 0 | 0
Enrolled Population
N=180
Non-inferiority criteria met for primary efficacy endpoint
*All-cause mortality
25% prespecified non-inferiority margin
Rate 7.8%
12.3%
Upper 1-sided 97.5% CI
P <0.0001
Primary Efficacy Endpoint at 1 Year*

## Slide 8
All-Cause Mortality at 2 Years
15.4%
All Cause Mortality %
Number at Risk
Number of Events
180
176
164
112
0
4
14
27
2.2%
7.8%

## Slide 9
All-Cause Mortality at 2 Years
15.4%
All Cause Mortality %
Number at Risk
Number of Events
180
176
164
112
0
4
14
27
2.2%
7.8%

## Slide 10
Cardiovascular Mortality at 2 Years
7.4%
Cardiovascular Mortality %
112
13
1.7%
6.2%
Number at Risk
Number of Events
180
176
164
0
3
11

## Slide 11
Cardiovascular Mortality at 2 Years
7.4%
Cardiovascular Mortality %
112
13
1.7%
6.2%
Number at Risk
Number of Events
180
176
164
0
3
11

## Slide 12: Hemodynamic Valve Performance
Chart: Mean Gradient (mm Hg)
| Baseline (n=180) | 30 Days (n=168) | 6 Months (n=153) | 1 Year (n=148) | 2 Years (n=113)
Mean Gradient (mmHg) | 8.7 | 3.9 | 4.3 | 4.2 | 4.2
EOA (cm ) | 2.7 | 2.9 | 2.7 | 2.8 | 2.8
2

## Slide 13: Paravalvular Regurgitation
Chart: 
| 30 Days (n=167) | 6 Months (n=153) | 1 Year (n=147) | 2 Years (n=113)
None/Trace | 0.833 | 0.942 | 0.918 | 0.956
Mild | 0.162 | 0.052 | 0.082 | 0.044
Moderate | 0.006 | 0.007 | 0 | 0

## Slide 14: NYHA Functional Class
Chart: 
| Baseline (n=180) | 30 Days (n=170) | 1 Year (n=148) | 2 Years (n=116)
Class I | 0 | 0.535 | 0.595 | 0.595
Class II | 0.322 | 0.365 | 0.324 | 0.302
Class III | 0.628 | 0.1 | 0.081 | 0.103
Class IV | 0.05 | 0 | 0 | 0

## Slide 15: The ALIGN AR Trial: 2 Year Follow-up Conclusions
Mortality was primarily noncardiac between 1 and 2 years in this population with numerous comorbidities
Hemodynamic valve performance was favorable beyond 1 year with average EOA >2.5 cm2 and mean transvalvular gradient <5 mmHg and no cases of hemodynamic valve deterioration
The favorable LV remodeling findings observed by echocardiography at 1 year were sustained at 2 years
The improvements in NYHA functional class and KCCQ-OS observed at 1 year were preserved at 2 year follow-up