## Slide 1: TAVR with JenaValve for Symptomatic Aortic Regurgitation in High Surgical Risk Patients
Outcomes in 500 patients from The ALIGN AR Trial
Raj R. Makkar, MD on behalf of ALIGN-AR investigators Associate Director, Cedars-Sinai Heart Institute

## Slide 2: Disclosure of Relevant Financial Relationships
Within the prior 24 months, I have had a relevant financial relationship with a company producing, marketing, selling, re-selling, or distributing healthcare products used by or on patients:
Nature of Financial Relationship | Ineligible Company
Grant/Research Support | Edwards Lifesciences, Medtronic, Abbott, Boston Scientific, Jenavalve
Consultant Fees/Honoraria | None
Individual Stock(s)/Stock Options | None
Royalties/Patent Beneficiary | None
Executive Role/Ownership Interest | None
Other Financial Benefit | None

## Slide 3: Background
Untreated severe, symptomatic AR is associated with high mortality especially in those with NYHA III/IV symptoms1
Surgery is the only recommended intervention for patients with native severe AR2
Yet, a significant proportion of patients with severe AR remain untreated with surgery3
Use of off-label TAVR devices for AR is associated with valve embolization (10%) and ≥moderate PVR (10%)4 both of which increase mortality5
Dujardin K, et al. Circulation. 1999;99:1851-1857.
Otto CM, et al. Circulation. 2021;143:e72-e227.
Thourani VH, et al. Structural Heart. 2021;5:608-618.
Poletti E, et al. PURPOSE. J Am Coll Cardiol Intv. 2024;1597–1606
Poletti E, et al. PANTHEON. J Am Coll Cardiol Intv 2023;16:1974–1985
Mortality >70% at 5 years
with NYHA III-IV Status

## Slide 4: Trilogy Porcine Pericardial Valve
27-31F
opening
Alignment
Aligns THV with native cusps
Positioning/Anchoring
Locators “clip” onto native leaflets forming a natural seal and stable securement
Deployment
Large open cells provide access to low coronaries. Flared sealing ring conforms to annulus
The Trilogy THV System is for Investigational Use Only in the United States and is Limited by Federal (or United States) law for this use.

## Slide 5: ALIGN AR Study Design
Clinical Evaluation, Echocardiography, Functional and QoL Assessment at Baseline, 30 Days, 6 Months, 1 Year and Annually up to 5 Years
Comparison with Prespecified Performance Goal
Multicenter, Non-blinded, Single Arm Evaluation of Patients with Symptomatic ≥3+ Aortic Regurgitation at High Risk for Surgery
TAVR with JenaValve Trilogy System
30 Day Primary Safety Endpoint
1 Year Primary Efficacy
Endpoint

## Slide 6: Key Inclusion and Exclusion Criteria
Inclusion
Adult patients with moderate-to-severe or severe (Grade ≥3) AR assessed according to ASE criteria
NYHA Class II or greater symptoms
High-risk for surgery defined by the heart team
Exclusion
Congenital unicuspid or bicuspid aortic valve
Ascending aorta diameter >5.0 cm
Previous prosthetic aortic valve
Mitral regurgitation > moderate
CAD requiring revascularization

## Slide 7: Top 10 Enrolling Sites
 | Principal Investigator | Number of Patients
Cedars-Sinai Medical Center | Raj Makkar | 96
Columbia University Medical Center/NYP | Torsten Vahl | 57
Piedmont Heart Institute | Vinod Thourani | 30
Rutgers/Robert Wood Johnson | Mark Russo | 27
University of Washington | Jamie McCabe | 27
California Pacific Medical Center | David Daniels | 25
University of Michigan | Stan Chetcuti | 24
Medstar/Washington Hospital Center | Lowell Satler | 22
Intermountain Health | Brian Whisenant | 19
Oregon Health Sciences University | Firas Zahr | 19
Scripps Health | Curtiss Stinis | 17
500 patients were enrolled at 28 US sites

## Slide 8: Baseline Characteristics (n=500)
Demographics and Co-Morbidities |  |  | Vascular & Other Co-Morbidities | 
Age (years) | 76.6 ± 10.3 |  | Atrial Fibrillation | 39.0%
Female | 46.2% |  | Pulmonary Hypertension | 19.8%
BMI – kg/m2 | 25.6 ± 5.7 |  | Prior Permanent Pacemaker | 15.2%
STS Score | 3.9 ± 3.3 |  | Left Bundle Branch Block | 7.6%
NYHA Class III or IV | 61.6% |  | RBBB/Bifascicular Block | 12.2%
Hypertension | 79.2% |  | Prior CABG | 9.2%
Diabetes | 16.0% |  | Prior PCI | 19.0%
Renal Insufficiency | 29.8% |  | Prior CVA | 9.6%
Right Ventricular Dysfunction | 5.2% |  | Carotid Disease | 8.0%
Prior Endocarditis | 5.8% |  | Peripheral Arterial Disease | 10.0%

## Slide 9: Screening and Patient Disposition
2 Year Follow-up* (n=206)
Missing or visit window not open (N= 104)
Lost to Follow-up (N=1) Patient Withdrew (N=3) Physician Withdrew (N=1) Other Reason Withdrew (N=2)
Missing or visit window not open (N= 282) Lost to Follow-up (N=1)
Patient Withdrew (N=3) Other (N=1)
6 Month Follow-up* (n=426)
Missing or visit window not open (N= 74)
1 Year Follow-up* (n=389)
*Available as of 3/22/2025
Patients Screened
(n=986)
Continued Access
Population
(n=320)
Pivotal Trial
Population
(n=180)
30 Day Follow-up Status
(n=500)
*Patients may have >1 exclusion
Patients Ineligible* (n=486)
Other Criteria (n=235)
CT Anatomic Criteria (N=183)
Perimeter (N=99)
Aortic Angle >70 (N=36)
Aorta >50 mm (N=13)
Aortic Length <55 mm (N=12)
Bicuspid (N=23)
ECHO Criteria (N=145)
AR Severity <3+ (N=120) LVEF <25% or Missing (N=10) MR >2+ (N=25)
Patients Enrolled (n=500)

## Slide 10: Index Procedure
Procedural Factors | 
Valve Size Implanted Small Medium Large | 24.8% 24.1% 51.0%
Post-Dilation | 3.7%
Procedure Time (min) | 69.5 ± 33.3
Sheath Time (min) | 28.2 ± 27.1
Procedural Complications | 
In-procedural Death | 0
Annular Rupture | 0
Ventricular Perforation | 0
Coronary Obstruction | 0
Valve Embolization | 1.6% (8)
Aortic Dissection | 0.6% (3)

## Slide 11: Procedural Outcomes
Technical Success | 95.2%
Procedural Death | 0.0%
Surgery or intervention related to device/procedure | 3.6%
THV deployment success | 98.0%
Delivery system success | 99.2%
Device Success | 96.4%
Procedural Death | 0.0%
THV deployment success | 98.0%
Patient prosthesis mismatch (EOAi ≤0.85) at 30 days | 0.4%
AOV gradient <20 mmHg at 30 days | 100.0%
AOV peak velocity <3 m/s at 30 days | 100.0%
Total AR Regurgitation: Moderate (n=3) or Severe (n=0) at 30 days | 0.6%

## Slide 12: Primary Safety Endpoint at 30 Days
ALIGN-AR (n=500) | 
Primary Safety Composite Endpoint | 26.2% (131)
All Cause Mortality | 1.4% (7)
Cardiovascular Mortality | 1.2% (6)
Any Stroke | 2.0% (10)
Disabling Stroke | 0.8% (4)
Non-disabling Stroke | 1.2% (6)
Major/Life Threatening Bleeding | 3.2% (16)
Major Vascular Complication | 2.8% (14)
Acute Kidney Injury Stage 2,3 or Dialysis (7 Days) | 0.6% (3)
Surgery/Intervention Related to the Device | 4.4% (22)
New Pacemaker Implantation | 23.3% (99)
Pre-existing Pacemaker | 15.2% (76)
≥ Moderate Total Regurgitation | 0.6% (3)

## Slide 13: Primary Safety Endpoint at 30 Days
ALIGN-AR (n=180) |  | ALIGN-AR CAP (n=320)
Primary Safety Composite Endpoint | 26.7% (48) | 25.9% (83)
All Cause Mortality | 2.2% (4) | 0.9% (3)
Cardiovascular Mortality | 1.7% (3) | 0.9% (3)
Any Stroke | 2.2% (4) | 1.9% (6)
Disabling Stroke | 1.1% (2) | 0.6% (2)
Non-disabling Stroke | 1.1% (2) | 1.3% (4)
Major/Life Threatening Bleeding | 4.4% (8) | 2.5% (8)
Major Vascular Complication | 3.9% (7) | 2.2% (7)
Acute Kidney Injury Stage 2,3 or Dialysis (7 Days) | 1.1% (2) | 0.3% (1)
Surgery/Intervention Related to the Device | 5.0% (9) | 4.1% (13)
New Pacemaker Implantation | 24.0% (36) | 23.0% (63)
Pre-existing Pacemaker | 16.7% (30) | 14.4% (46)
≥ Moderate Total Regurgitation | 0.6% (1) | 0.7% (2)

## Slide 14: Primary Safety Endpoint at 30 Days*
10%
25%
40%
p<0.0001
40.5% prespecified non-inferiority margin
30.1%
Upper 1-sided 97.5% CI
Rate
26.2%
Feldman TE, Reardon MJ, Rajagopal V, et al. The REPRISE III Randomized Clinical Trial. JAMA. 2018; 319: 27-37.
Makkar RR, Cheng W, Waksman R, et al. Self-expanding intra-annular versus commercially available transcatheter heart valves in high and extreme risk patients with severe aortic stenosis (PORTICO IDE): a randomised, controlled, non- inferiority trial. Lancet. 2020; 369: 669-683.
Thiele H, Kurz T, Feistritzer H-J, et al. Comparison of newer generation self-expandable vs. balloon-expandable valves in transcatheter aortic valve implantation: the randomized SOLVE-TAVI trial. Eur Heart J. 2020; 41:1890-1899.
*Composite of 30-day all-cause mortality, any stroke, life-threatening/major bleeding, major vascular complication,
AKI ≥2 or dialysis, valve intervention, new permanent pacemaker, ≥ moderate valvular regurgitation

## Slide 15: All Cause Mortality

## Slide 16: Cardiac Mortality

## Slide 17: Primary Efficacy Endpoint at 1 Year*
10%
20%
30%
 | 25% prespecified non- inferiority margin |  | 
Rate Upper 1-sided 8.1% 97.5% CI 10.7% |  |  | 
p<0.0001
Fiedler AG, Bhambhani V, Laikhter E, et al. Heart. 2018; 104:835-840.
Kamath AR, Varadarajan P, Turk R, et al. Circulation. 2009: 120:S134-8.
Turk R, Varadarajan P, Kamath A, et al. Ann Thorac Surg. 2010; 89: 731-737.
Sampat U, Varadarajan P, Turk R, Kamath A, Khandhar S, Pai RG. J Am Coll Cardiol. 2009; 54:452-457.
*All Cause Mortality at 1 Year

## Slide 18: All Stroke

## Slide 19: Hemodynamic Valve Performance
7.5
4.3
2.7
3.0
2.8
2.8
2.8
1.5
2.0
2.5
3.0
3.0
5.0
7.0
9.0
EOA (cm2)
Mean Gradient (mm Hg)
Mean Gradient (mmHg)
EOA (cm2)
Baseline
n=496
6 Months
n=388
4.2
1 Year
n=288
4.2
2 Years
n=116
3.7
30 Days
n=474

## Slide 20: Paravalvular Regurgitation
82.6%
93.6% 93.4%
95.7%
6.2% 6.3% 4.3%
0.3%
0.4% 0.3%
17.0%
20%
0%
40%
100%
30 Days
6 Months
1 Year
2 Years
80% Moderate
60%
Mild
None/Trace

## Slide 21: Left Ventricular Remodeling
38.4
36.4
32.3
30.8
30.5
25
30
35
40
LV ESV Index
Milliliters/m2
82.2
71.2
64.0
63.1
65.0
85
80
75
70
65
60
55
LV EDV Index
Milliliters/m2
p<0.0001
p<0.0001
152.3
117.0
112.0
105.5
114.3
80
100
120
140
160
LV Mass Index
Grams/m2
p<0.0001

## Slide 22: NYHA Functional Class
53%
62%
59%
38%
39%
31%
31%
58%
9%
8%
10%
4%
0%
20%
40%
60%
80%
100%
Baseline (n=500)
2 Years (n=125)
30 Days (n=478)
Class I Class II
1 Year (n=331)
Class III Class IV

## Slide 23: KCCQ-Overall Summary
KCCQ-OS scores range from 0 to 100, with higher scores indicating better health status. A large change is defined as ≥20 points compared with baseline.
100 |  | ∆19.4 ± 25.5; p<0.0001 |  | 100% 80% 60% 40% 20% 0% | Alive and Well (KCCQ ≥60 Without Decrease of >10) 84.0% 79.8% 1 Year 2 Years
 |  | ∆21.5 ± 22.6; p<0.0001 |  |  | 
80 |  |  |  |  | 
60 |  |  |  |  | 
40 |  | 81.1 | 79.1 |  | 
20 |  | 58.5 |  |  | 
0 |  |  |  |  | 
Baseline (N=497) |  |  |  |  | 

## Slide 24: Pacemaker by Valve Size*
16.8%
18.3%
28.9%
0%
10%
30%
20%
40%
23mm 25mm 27mm
n=101 n=104 n=211
*Excludes 76 patients with prior pacemaker and 8 patients who did not receive the Trilogy THV

## Slide 25: Predictors of New Pacemaker
 | Odds Ratio | 95% Wald Confidence Limits |  | p-value
History of Congestive Heart Failure | 1.92 | 1.15 | 3.21 | 0.01
Baseline Severe AR vs. Moderate to Severe | 1.88 | 1.10 | 3.19 | 0.02
Annular Perimeter ≥85mm | 2.08 | 1.24 | 3.46 | 0.005
Baseline RBBB | 6.66 | 3.48 | 12.75 | <0.0001
Variables not associated with new pacemaker included NYHA classification, STS score, LBBB, Oversizing, Depth by echocardiography, LVEF, LV Dimensions, LV Volume, LV Mass, Site Experience, Enrollment Tercile, Baseline KCCQ

## Slide 26: Conclusions
In the largest cohort of high-risk patients with symptomatic native aortic regurgitation undergoing TAVR with the JenaValve:
The primary safety and efficacy outcomes achieved prespecified
performance goals
  Safety - 26.2% vs. 40.5%, p<0.0001
  Efficacy (all-cause mortality) - 8.1% vs. 25%, p<0.0001
High device success rates (96.4%) and acceptable rates of complications highlight a favorable risk-benefit profile
  Procedural death 0%, 30-day death 1.4%, valve embolization 1.6%, 30-day disabling stroke 0.8%
Favorable valve performance was evidenced by consistently low valve
gradients, large valve areas, and low rates of valvular regurgitation
  EOA ~ 2.8 cm2 and mean gradient 4.2 mm Hg, ≥ moderate total AR 0.9% at 2 years

## Slide 27: Conclusions (continued)
Sustained improvement in functional status (NYHA class improvement) and patient-reported quality-of-life measures (KCCQ score improvement)
  NYHA class I/II: 90% at 2 years
  KCCQ-OS score 58.5 at baseline to 79.1 at 2 years
Significant reductions in LV end systolic and diastolic volumes and
regression of LV mass
  LVESVI 38.4 ml/m2 at baseline decreased to 30.5 ml/m2 at 2 years
  LV mass index decreased from 152.3 g/m2 to 114.3 g/m2
New pacemaker implantation rates are 23% reflecting underlying pathology
  Risk factors include larger annular sizes, heart failure, severity of AR, and RBBB

## Slide 28: The ALIGN AR Trial Clinical Implications
The TRILOGY THV is a unique TAVR device for symptomatic patients with ≥3+ AR. Upon FDA approval, such patients at high-risk for surgery will have a new and much-needed therapeutic option.
The technical success, reassuring safety profile and positive clinical outcomes in ALIGN-AR support a randomized trial comparing this device to SAVR in a patient population not at high-risk for surgery.
The Trilogy THV System is for Investigational Use Only in the United States and is Limited by Federal (or United States) law for this use.

## Slide 29: The ARTIST Trial Aortic Regurgitation Trial Investigating Surgery and Trilogy
Randomize 1:1
N=1016
Patient with 3-4+ Native Valve AR Requiring
AVR Based On Clinical Evaluation
Primary Noninferiority Endpoint at 12 months: 1) Death 2) Any symptomatic stroke 3) Urgent cardiac rehospitalizations with 10 Year follow-up
TRILOGY
THV TAVR
SAVR
Heart Team Deems Patient Suitable for
Screen Fail
No
No
TAVR and SAVR
Yes
Imaging Core Laboratory Confirms AR Severity and Case Review Board Confirms Indication and Suitability for Randomization
Yes
Study Chair
Martin Leon MD, Columbia
Principal Investigators
US
Raj Makkar MD, Cedars Sinai Vinod Thourani MD, Piedmont Heart Torsten Vahl MD, Columbia
OUS
Stephan Baldus, University Cologne Hendrik Treede, University Mainz